## Practical Data & Tools
Energy Efficiency Technician Practical Modules Requirements
|  | Equipment specifications: |  | Team toolbox; 5 learners 1 trainer
# | Equipment | Type | Specifications | Quantity of equipmet
1 | Energy logger | Energy / Power Logger | Temporary energy meter Battery powered (at least 2 weeks on full measurable load) Remote data capture with 1 year data comms included IP54 rated minimum Readable screen for settings and verification of measurements on site Settable logging intervals with 5 minute as minimum - Instantaneous or averaged or total per interval as per measured sensor requirement Utilising CT's as per above spec Fused and insolated clamp-on voltage connectors - 420V rated - 600V insulation - 1A fused All associated leads, cables and probes included. Protective pouch or storage means Operations manual included | 2
2 | Power analyser | Hand-held power meter | Clamp-on with safety probes CAT IV 600V; CAT III 1000V insulation True RMS measurement Single or three-phase capable 50 Hz At least 3% accuracy - 0 - 1000 Amps AC - 0 - 600V AC - Power Factor - kW - kVA - Others bonus All associated leads and probes included Protective pouch or storage means Operations manual included | 2
3 | Light meter | Hand-held light intencity meter | Capable reading incandescent, Fluourescent, HID, LED light types Colour correction filter selction 0 - 200 000LUX +- 6% accuracy Self-calibration capable All associated leads, cables and probes included. Protective pouch or storage means Operations manual included | 2
4 | Occupancy logger* | Occupancy Logger | Capable to log occupation of an area Battery opperated Detection range 5m minimum Detecton angle 90 degrees minimum horizontal and vertical All associated leads, cables and probes included. Protective pouch or storage means Operations manual included | 6
5 | Ballast discriminator* | Ballast Discriminator | Capable of indirectly provide feedback on the type of control gear used by fluourescent tubes and other discharge type lighting systems. Most operate on blinking rate of light source All associated leads, cables and probes included. Protective pouch or storage means Operations manual included | 2
6 | Data loggers | Multi Logger | Battery powered (at least 2 weeks on full measurable load) Remote data capture with 1 year data comms included IP54 rated minimum Readable screen for settings and verification of measurements on site Settable logging intervals with 5 minute as minimum - Instantaneous or averaged or total per interval as per measured sensor requirement Capable of reading multiple sensors - A ac TRMS (0-5A, 0-333mV) x 3 - Temperature (Thermalcouple J or K-type) x 2 - Pulse counting x 1 - 4-20mA x 1 - 0-5V as bonus x 1 All associated leads, cables and probes included. Protective pouch or storage means Operations manual included | 2
7 |  | Electricity Amp logger | Battery powered (at least 2 weeks on full measurable load) Remote data capture with 1 year data comms included IP 54 rated minimum 4 x Amp inputs - proprietary or universal signal type, TRMS Settable logging intervals with 5 minute as minimum - Instantaneous, averaged or cumulative as per sensor and measurement requirement All associated leads, cables and probes included. Protective pouch or storage means Operations manual included | 6
8 |  | Split Core CTs for multi or electricity Amp logger | Class 2 accuracy 600V Insulation Durably hinged with catch, no separate screws to connect split cores. 0-30A x 9 (Capable to fit in sub DB) 0-100A x 6 (Capable to fit in sub DB) 0-250A x 3 0-500A x 3 Last two types can also be replaced by flexi-ratio CT x 3 Closely similar Amp sizes acceptable. | 2
9 |  | Temperature probes for multi-logger | Capable to be used with suggested logger -40 - 120DegC 0.5DegC accuracy 3m wire Durable contact able to be clamped to piping or heated surface or submersible to 1m | 2
10 |  | Pulse Leads for multi-logger | 3m, 2-wire pulse leads to connect to related pulse output Twisted pair and electro-magnetically sheilded Low resistance material, 0.5mm^2 wire thickness Connection capable to multilogger | 2
11 | Temperature thermometer, thermal camera | Hand-held thermal IR measurement | LCD display Battery operated Measuring range -35DegC - 500DegC IR and thermocouple capable +-2% of reading or +-2DegC, whichever is greater Field of view 12:1 Adjustible emissivity Laser marker All associated leads, cables and probes included. Protective pouch or storage means Operations manual included | 2
12 |  | Hand-held IR Camera | 4-hour battery life Spectral range 8-14μm IP54 rating Measurement range -20DegC - 300DegC minimum Accuracy +-4DegC All associated leads, cables and probes included. Protective pouch or storage means Operations manual included | 2
13 | Wet and dry bulb thermomenter (or psychormeter) | Wetbulb temperature and humidity | Handheld battery powered -10 DegC - 60 DegC 0-100% Relative Humidity 1 DegC resolution +-1 DegC accuracy 5% RH resolution +-5% RH Accuracy All associated leads, cables and probes included. Protective pouch or storage means Operations manual included | 2
14 | Flow rate meter for water and air/flow hood meter | Handheld Air Flow Meter | Battery operated 0.4 - 40 m/s -10DegC - 50DegC operational All associated leads, cables and probes included. Protective pouch or storage means Operations manual included | 2
15 |  | Ultrasonic flow meter | Portable and battery operated for 4 hours with means for mains to be added for longer periods of measure IP54 rated On-site setup through interactive screen 20mm - 500mm diameter capability at least Steel, Stainless Steel, copper, PVC, Galvanised, Alluminium, Cast Iron. With or without liner. Clamp-on probes with suitable commonly available accoustic gel Multi reflex setup possible Water, oil, glycol, diesel, liquid LPG capable at least 0 - 7m/s minimum +-1% accuracy -30DegC - 160DegC Self logging capability and/or pulse/4-20mA output to multilogger - Settable intervals - Flow speed, Mass flow, Volume flow - Average or cumulative per interval Pipe thickness guage included Suitable means to be left on site for 1 week All associated leads, cables and probes included. Protective pouch or storage means Operations manual included | 2
16 | Ultrasonic leak detector | Ultrasonic leak detector | Hand-held, battery operated capable of 2 hours operation minimal Capable detecting water, air and gas leaks Operational between -10DegC and 50DegC ambient All associated leads, cables and probes included. Protective pouch or storage means Operations manual included | 2
17 | Laser distance measurement (50m range) | Laser distance meter | 0.05m - 60m minimum Battery operated Length, Square Area and volume capable 0-reference changeable Addition of successive measurements All associated leads, cables and probes included. Protective pouch or storage means Operations manual included | 2
18 |  | Toolbox | Big Jim type, lockable, durable, portable - Full set of electrician's screwdrivers - Large side cutters isolated - Long nose pliers isolated - Medium flat nose pliers isolated - Packet of 300mm cable ties (50 minimum) - Nitto type isolation tape x 4 (Black, Red, Yellow, Blue) - Measuring tape 10m - Universal DB key - Electrician gloves (600V isolation, elbow length) - Cleaning Rags - Sportman stopwatch - LED pencil flashlight - Laminated list of toolbox contents with tick-place for content check | 1
19 |  | Stepladder | Foldable, 4-step | 1